## Slide 1
Basic Traffic test with xGenius
Guide & Slides corresponding to software ver. 2.3.13
Ethernet Traffic test

## Slide 2: Zeus & xGenius
Zeus & xGenius provide deep insights to design, install, maintain, troubleshoot and engineer communications resources of Mobile Operators, Power Utilities, Finace, Labs and R+D centers. The unit is able to verify Ethernet/IP, SyncE, PTP, GbE, IRIG-B, 1PPS, T1/E1, G703, Serial Datacom, C37.94, GOOSE, SV, MMS, Round-trip & One-way Delay tests at all interfaces.
It also has a set of programmable filters to capture live data traffic at wire-speed

## Slide 3: Interfaces & time References
xGenius generates and analyses different kinds of 1544 kb/s and 2048 kb/s signals and also unframed 1544 / 2048 kb/s signals.
PHM
RJ45-C
BNC-C
SFP+ A
SFP+ B
RJ45-A
RJ45-B
SFP-A
Micro SD
SMB-2
SMB-1
SMA
SMB-C
RJ45
SFP-B
 | GNSS | Eth/IP | PTP | NTP | ToD | IRIG-B | PPS | PP2S | SyncE | T1/E1 | MHz | Data
RJ45-A |  | in/out | in/out | in/out |  |  |  |  | in/out |  |  | 
RJ45-B |  | in/out |  |  |  |  |  |  | in/out |  |  | 
SFP-A |  | in/out | in/out | in/out |  |  |  |  | in/out |  |  | 
SFP-B |  | in/out |  |  |  |  |  |  | in/out |  |  | 
SFP+A |  | in/out | in/out | in/out |  |  |  |  | in/out |  |  | 
SFP+B |  | in/out |  |  |  |  |  |  | in/out |  |  | 
BNC-C |  |  |  |  |  |  |  |  |  | in/out | in/out | 
RJ45-C |  |  |  |  | in |  | in | in |  | in/out | in/out | 
SMB-C |  |  |  |  |  |  |  |  |  |  |  | 
SMA | in |  |  |  |  |  |  |  |  |  |  | 
SMB-1 |  |  |  |  |  | in | in | in |  |  |  | 
SMB-2 |  |  |  |  |  | out | out | out |  |  |  | 
RJ-45 |  |  |  |  | in/out | in/out | in/out | in/out |  | in/out | in/out | 
PHM |  |  |  |  |  |  |  |  |  |  |  | in/out

## Slide 4: The basic 802.3 MAC frame format
Ethernet has adopted many of the transmission media available, including coaxial, UTP, STP, multimode and monomode fibers, in order to support the various market demands.

## Slide 5: The Traffic test
The test described aims to teach how to set up the tester to generate and receive real traffic with specific frames, bit rates, and events. With a fea chabves it will be possible to evaluate other traffic parameters such as throughput, latency, frame loss rate, etc. The document also describes specific formats for tracing and reporting the results of this test.
Routed Network
1
ALBEDO
xGenius
2
ALBEDO
xGenius
Routed Network
1000BASE-T
1
ALBEDO
xGenius
PHY Loop
Port A
Port B
Port A
Using 2 testers
Test 1
Using 1 tester
Test 2

## Slide 6: Connecting the BNC cables to Port C
xGenius Ports A and B have the connector for the RJ45 cables
xGenius also support optical media and higher bit rates but we are going to execute this test in 1000BASE-T
Connect the Cable to Port A in tester A
Connect the Cabel to Port B in tester B
1
1

## Slide 7: Configure the tester
Restore defaults: Home > [System] > Reset to factory settings := Yes
Set operation mode: Home > Test > Mode := Ethernet endpoint
2
1

## Slide 8: Configure the PHY layer Tester 1
Configure in Tester 1 the Physical layer:
Config > Port A > Physical Layer > Connector := Electrical
Config > Port A > Port Mode := Tx/Rx
In the remote site configure also Tester 2* and Port B to loop back the frames that will be sent in Tester 1:
Config > Port B > Port Mode := LoopBack
(*) If you want just to learn how to program you can use Port A and B of the same tester
3
1

## Slide 9: Configure the PHY layer Tester 2
Configure in Tester 2 the traffic profile:
Config > Port B > Physical Layer > Connector := Electrical
Config > Port B > Port Mode := LoopBack
(*) If you want just to learn how to program you can use Port A and B of the same tester
3
1

## Slide 10: Configure Traffic Profile
The tester can generate up to 8 traffic flows simultaneously, set up the 1st one:
Config > Port A > Bandwidth profile > Flow A1 := Ramp
Configure the bandwidth parameters:
Config > Port A > Bandwidth profile > Flow A1 > Tramission rate := [Rate fr/s | Rate (Mb/s) | Rate (%) …]
You can also program a payload for instace a PRBS: Config > Port A > Payload type := BERT Config > Port A > BERT patterns := ITU PBS 2^11
2
1

## Slide 11: Press RUN
You can spot the frame and the test profile configured at Filter A1 including the type of traffic to be generated at Port A including the flow, the Physical layer, the Ethernet frame, MAC addresses, the PRBS generated, etc.
Press RUN to activate the traffic generation
2
1

## Slide 12: Press RUN
You can see now if the connection was fine looking the autonegotiation:
Results > Port A > Autonegotiation
Press LED symbol
Immediately the Rx and Tx LED will turn green indicating the generation and the reception of traffic
As you have selected a PRBT spot the LED indicating that the pattern is being received correctly
You can spot other results for instance some statistics regarding traffic: Results > Port A > Bandwidth statistics > General statistics
3
2
4
1
5

## Slide 13: Insert events in Tester 1
This is a good practice to understand how it all works
Test > Insertion > Insert to := Port A
Test > Insertion > Event to insert := TSE
Test > Insertion > Mode := Single
Press Event to insert one TSE error
Spot how the error appears in the LED and the tester shows an E label
2
5
4

## Slide 14: Event logger enable
Global counts, statistics and LEDs provide information about which events and how many of them have been registered but they do not say too much about how they are distributed in time. These information is supplied at the Event Logger a graphical representation tool that permits to trace the evolution of key parameters. The first thing to to is to enable it:
Home > Test > Event logger setup > Enable := Yes
Clear the currently selected filters: Clear all filters := Yes
1

## Slide 15: Select the Events to Log
Now you need to select which events you want to trace are categorized in classes:
Test > Event Logger setup > [ Port A | Port B | System event | Debug events]
You will find out several groups of events and statistics: Port A event filter > Port A event filter := [Anom/defts | Bwd | Frame | SLA | BERT]
For our particular test select a bundle or events in Port A : Port A event filter > Bandwidth stats > Filter A1 := [Ethernet | IPv4 | IPv6 | UDP]
1
3
2
2
1
3

## Slide 16: Display the logs
With the Event Logger you can select one or various events and trace them so that all changes along with the time and date these changes are registered are recorded with a 1 second resolution. The event logger provides representations and zoom levels to enable event analysis at different time scales:
Go to Results > Event logger
To select a saved trace double click at (…) to display the folder of traces
Select and Open one of the files
1
2
3
3

## Slide 17: Display chronogram & histogram
To display the logger open the actual trace always in red
Select the event you want to display in the chronograme
You can see the value evolution
To know the exact value at one specific moment press on the screen
1
3
4
4

## Slide 18: Scale the results on screen
You can change the scale to see in more details the results
Using the auxiliary bars change the vertical scale
Reset to the original scale
2
3
1

## Slide 19: Exporting Logs
To export log files in CSV format:
Go Home > [File icon] > Configuration files | Report files > Internal memory to list the files currently stored in the unit
Select the files you want to export: File name > Export
Choose an external device: micro SD | USB
If no external device a popup will appear: ‘No devices present’
2
3
1
1

## Slide 20
That’s all
www.albedotelecom.com